## Flexisource
| TG-43 CONSENSUS |  |  |  |  | Isodose Control HDR 192Ir Flexisource
Dose calculation for photon-emitting brachytherapy sources with average energy higher than 50 keV: Full Report of the AAPM and ESTRO
Med. Phys. 39 (2012) 2904-2929
| L = | 1.1125 |  | cGy/(h U)
Interpolated / extrapolated data are boldface / underlined. Values inside the source are in italics.
| gL(r) |  |  |  | F(r,q) | [9 empty] | QA Along-away
| L = | 0.35 | cm |  |  |  | Distance from Active Source Center (cm)
| r (cm) | gL(r) L = 0,35 cm unbounded |  | Theta/deg | 0 | 0.25 | 0.5 | 1 | 2 | 3 | 4 | 5 | 7.5 | 10 | z / cm | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
| 0 | 0.9911 |  | 0 | 0.67195 | 0.67195 | 0.6535 | 0.6166 | 0.6256 | 0.6469 | 0.6715 | 0.6947 | 0.7382 | 0.7741 | 7 | 0.016418 | 0.016716 | 0.01701 | 0.017303 | 0.01758 | 0.017882 | 0.017863 | 0.01705 | 0.015586 | 0.013845 | 0.012044 | 0.01042
| 0.25 | 0.9911 |  | 1 | 0.67075 | 0.67075 | 0.6522 | 0.6151 | 0.6291 | 0.6518 | 0.6775 | 0.6988 | 0.7436 | 0.7767 | 6 | 0.022052 | 0.022575 | 0.023068 | 0.023618 | 0.024081 | 0.024427 | 0.024238 | 0.022503 | 0.019952 | 0.017134 | 0.014507 | 0.012259
| 0.5 | 0.9965 |  | 2 | 0.66905 | 0.66905 | 0.6511 | 0.6152 | 0.6378 | 0.664 | 0.6876 | 0.7111 | 0.7514 | 0.7826 | 5 | 0.031225 | 0.032198 | 0.033172 | 0.034192 | 0.034839 | 0.035141 | 0.034324 | 0.030501 | 0.025817 | 0.021294 | 0.0175 | 0.014315
| 0.75 | 0.9983 |  | 3 | 0.66285 | 0.66285 | 0.6516 | 0.6291 | 0.6504 | 0.6772 | 0.6991 | 0.7192 | 0.7587 | 0.7892 | 4 | 0.047414 | 0.049572 | 0.051776 | 0.053651 | 0.054486 | 0.054095 | 0.05111 | 0.042559 | 0.033772 | 0.026506 | 0.020819 | 0.016543
| 1 | 1 |  | 5 | 0.6714 | 0.6714 | 0.6652 | 0.6528 | 0.6758 | 0.6982 | 0.7185 | 0.7369 | 0.7745 | 0.8024 | 3 | 0.081463 | 0.087027 | 0.092656 | 0.095654 | 0.096086 | 0.090719 | 0.081175 | 0.060534 | 0.044061 | 0.032433 | 0.024393 | 0.01875
| 1.5 | 1.0017 |  | 7 | 0.69425 | 0.69425 | 0.6901 | 0.6818 | 0.703 | 0.7246 | 0.7428 | 0.7603 | 0.7923 | 0.816 | 2 | 0.17777 | 0.19721 | 0.21206 | 0.2129 | 0.20298 | 0.17016 | 0.13612 | 0.0854 | 0.055843 | 0.038445 | 0.027676 | 0.020688
| 2 | 1.0037 |  | 10 | 0.73455 | 0.73455 | 0.7312 | 0.7245 | 0.7442 | 0.7625 | 0.7801 | 0.7941 | 0.8207 | 0.8407 | 1.5 | 0.31502 | 0.35861 | 0.37909 | 0.3599 | 0.32271 | 0.24141 | 0.17652 | 0.099472 | 0.061498 | 0.04105 | 0.029047 | 0.021474
| 3 | 1.0051 |  | 12 | 0.7618 | 0.7618 | 0.7599 | 0.7561 | 0.7701 | 0.785 | 0.799 | 0.812 | 0.8353 | 0.8543 | 1 | 0.7148 | 0.84801 | 0.81233 | 0.67976 | 0.54213 | 0.33989 | 0.22313 | 0.11249 | 0.066212 | 0.043138 | 0.030113 | 0.022053
| 4 | 1.0034 |  | 15 | 0.80265 | 0.80265 | 0.7989 | 0.7914 | 0.8036 | 0.8165 | 0.8287 | 0.8394 | 0.8571 | 0.8727 | 0.5 | 3.3359 | 3.4456 | 2.1958 | 1.3542 | 0.88584 | 0.44674 | 0.26416 | 0.12195 | 0.069403 | 0.044477 | 0.030748 | 0.022388
| 5 | 0.9987 |  | 20 | 0.8521 | 0.8521 | 0.8496 | 0.8446 | 0.851 | 0.8609 | 0.8703 | 0.8777 | 0.8892 | 0.8984 | 0 | 384888243.9808 | 15.5475 | 4.3088 | 1.9594 | 1.1125 | 0.49805 | 0.28126 | 0.12536 | 0.070429 | 0.044874 | 0.030932 | 0.022487
| 6 | 0.9912 |  | 25 | 0.89185 | 0.89185 | 0.8872 | 0.8779 | 0.8862 | 0.893 | 0.8988 | 0.9039 | 0.9119 | 0.9201 | -0.5 | 2.3081 | 3.45 | 2.1964 | 1.357 | 0.8852 | 0.44632 | 0.26394 | 0.12205 | 0.069573 | 0.044548 | 0.030766 | 0.02238
| 8 | 0.968 |  | 30 | 0.9167 | 0.9167 | 0.9126 | 0.9044 | 0.9108 | 0.9169 | 0.9207 | 0.9222 | 0.9323 | 0.9355 | -1 | 0.54799 | 0.84968 | 0.81488 | 0.68152 | 0.5427 | 0.34017 | 0.22278 | 0.11226 | 0.066268 | 0.043187 | 0.030091 | 0.022011
| 10 | 0.9349 |  | 35 | 0.93575 | 0.93575 | 0.9332 | 0.9281 | 0.9318 | 0.9364 | 0.941 | 0.9429 | 0.9488 | 0.9525 | -1.5 | 0.25018 | 0.35709 | 0.37963 | 0.36068 | 0.32346 | 0.24158 | 0.17665 | 0.099355 | 0.061338 | 0.041066 | 0.02907 | 0.021445
|  |  |  | 40 | 0.9549 | 0.9549 | 0.9511 | 0.9435 | 0.9477 | 0.9512 | 0.953 | 0.955 | 0.9583 | 0.9606 | -2 | 0.146 | 0.1945 | 0.21248 | 0.21323 | 0.20331 | 0.1704 | 0.13628 | 0.085416 | 0.055853 | 0.038395 | 0.027661 | 0.020698
|  |  |  | 45 | 0.9643 | 0.9643 | 0.962 | 0.9574 | 0.9604 | 0.9637 | 0.9666 | 0.9677 | 0.9666 | 0.9703 | -3 | 0.069928 | 0.084167 | 0.092142 | 0.095884 | 0.096206 | 0.090864 | 0.081266 | 0.060526 | 0.044028 | 0.0324 | 0.024352 | 0.018717
|  |  |  | 50 | 0.97335 | 0.97335 | 0.9718 | 0.9687 | 0.9713 | 0.9731 | 0.9754 | 0.9783 | 0.9784 | 0.98 | -4 | 0.041751 | 0.04726 | 0.051129 | 0.05351 | 0.054637 | 0.054141 | 0.051188 | 0.042618 | 0.033748 | 0.02648 | 0.020805 | 0.016511
|  |  |  | 55 | 0.9861 | 0.9861 | 0.9787 | 0.9746 | 0.9785 | 0.9805 | 0.9829 | 0.9831 | 0.9834 | 0.9859 | -5 | 0.028074 | 0.030513 | 0.03251 | 0.033904 | 0.034802 | 0.035163 | 0.03434 | 0.030535 | 0.025863 | 0.021284 | 0.017484 | 0.014306
|  |  |  | 60 | 0.9899 | 0.9899 | 0.9841 | 0.9822 | 0.9848 | 0.9873 | 0.9898 | 0.9903 | 0.9894 | 0.989 | -6 | 0.020033 | 0.021286 | 0.022511 | 0.023349 | 0.023958 | 0.024431 | 0.02425 | 0.022527 | 0.019914 | 0.017171 | 0.0145 | 0.012247
|  |  |  | 65 | 0.9929 | 0.9929 | 0.9892 | 0.9883 | 0.9904 | 0.9927 | 0.9937 | 0.9944 | 0.9945 | 0.9928 | -7 | 0.015062 | 0.015777 | 0.016515 | 0.01706 | 0.017447 | 0.01786 | 0.017866 | 0.017068 | 0.015586 | 0.013819 | 0.012066 | 0.010415
|  |  |  | 70 | 0.9961 | 0.9961 | 0.9933 | 0.9926 | 0.9937 | 0.9962 | 0.997 | 0.998 | 0.9954 | 0.9964
|  |  |  | 75 | 0.9974 | 0.9974 | 0.9951 | 0.9968 | 0.9964 | 0.9977 | 0.999 | 0.999 | 0.9994 | 0.9987
|  |  |  | 80 | 0.9992 | 0.9992 | 0.9995 | 0.9953 | 0.9996 | 0.9998 | 1.0004 | 1.0009 | 1.0021 | 1.0012
|  |  |  | 85 | 0.9995 | 0.9995 | 1.0004 | 0.9977 | 0.9989 | 0.999 | 1.0013 | 1.0013 | 1.0011 | 1.0013
|  |  |  | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
|  |  |  | 95 | 0.9997 | 0.9997 | 0.9991 | 0.9953 | 0.9989 | 1.0013 | 1.0028 | 1.003 | 1.0001 | 1.0015
|  |  |  | 100 | 0.9989 | 0.9989 | 0.9975 | 0.9954 | 0.9981 | 1.0003 | 1.0042 | 1.002 | 0.9986 | 0.9993
|  |  |  | 105 | 0.9983 | 0.9983 | 0.9964 | 0.9931 | 0.9957 | 0.9973 | 0.999 | 1.0008 | 1 | 1.0019
|  |  |  | 110 | 0.9958 | 0.9958 | 0.9932 | 0.9923 | 0.9931 | 0.9934 | 0.9937 | 0.9957 | 0.995 | 0.9937
|  |  |  | 115 | 0.9937 | 0.9937 | 0.9905 | 0.9863 | 0.9886 | 0.991 | 0.9948 | 0.9948 | 0.9918 | 0.9943
|  |  |  | 120 | 0.9898 | 0.9898 | 0.986 | 0.9846 | 0.9842 | 0.9848 | 0.9883 | 0.9896 | 0.9873 | 0.9878
|  |  |  | 125 | 0.9854 | 0.9854 | 0.9794 | 0.9768 | 0.9793 | 0.9819 | 0.9835 | 0.9829 | 0.9832 | 0.9847
|  |  |  | 130 | 0.97245 | 0.97245 | 0.971 | 0.9681 | 0.9709 | 0.9735 | 0.9757 | 0.9767 | 0.9775 | 0.9789
|  |  |  | 135 | 0.9629 | 0.9629 | 0.9615 | 0.9587 | 0.961 | 0.9649 | 0.9667 | 0.9669 | 0.9662 | 0.9698
|  |  |  | 140 | 0.9524 | 0.9524 | 0.95 | 0.9452 | 0.9489 | 0.9521 | 0.9552 | 0.9587 | 0.9606 | 0.9625
|  |  |  | 145 | 0.9366 | 0.9366 | 0.9349 | 0.9315 | 0.9333 | 0.9378 | 0.9415 | 0.9428 | 0.9461 | 0.9509
|  |  |  | 150 | 0.91815 | 0.91815 | 0.9146 | 0.9075 | 0.9137 | 0.9186 | 0.9222 | 0.9247 | 0.9318 | 0.937
|  |  |  | 155 | 0.891 | 0.891 | 0.8876 | 0.8808 | 0.887 | 0.8945 | 0.8999 | 0.905 | 0.9131 | 0.9188
|  |  |  | 160 | 0.8385 | 0.8385 | 0.8405 | 0.8445 | 0.8525 | 0.8613 | 0.8713 | 0.8776 | 0.8901 | 0.898
|  |  |  | 165 | 0.7834 | 0.7834 | 0.7866 | 0.793 | 0.8058 | 0.8188 | 0.831 | 0.8403 | 0.857 | 0.874
|  |  |  | 168 | 0.7479 | 0.7479 | 0.7511 | 0.7575 | 0.7703 | 0.786 | 0.8019 | 0.8123 | 0.834 | 0.8545
|  |  |  | 170 | 0.71078 | 0.71078 | 0.71515 | 0.7239 | 0.7414 | 0.7596 | 0.7758 | 0.7912 | 0.8175 | 0.8377
|  |  |  | 173 | 0.65853 | 0.65853 | 0.66345 | 0.6733 | 0.693 | 0.7147 | 0.7333 | 0.7499 | 0.785 | 0.8094
|  |  |  | 175 | 0.61415 | 0.61415 | 0.6196 | 0.6305 | 0.6523 | 0.6784 | 0.7007 | 0.7195 | 0.7603 | 0.7914
|  |  |  | 177 | 0.54197 | 0.54197 | 0.55005 | 0.5662 | 0.5985 | 0.6312 | 0.6601 | 0.6839 | 0.729 | 0.7659
|  |  |  | 178 | 0.47373 | 0.47373 | 0.48655 | 0.5122 | 0.5635 | 0.5992 | 0.6323 | 0.6591 | 0.7115 | 0.751
|  |  |  | 179 | 0.43988 | 0.43988 | 0.45335 | 0.4803 | 0.5342 | 0.5708 | 0.6057 | 0.6351 | 0.6929 | 0.7343
|  |  |  | 180 | 0.44188 | 0.44188 | 0.45215 | 0.4727 | 0.5138 | 0.5553 | 0.5913 | 0.6246 | 0.6804 | 0.7222