## General information
General Information
Source Information
Source name: | LDR wire
Radionuclide: | Ir-192
Model Number: | Ir-192 LDR wire 0.3 mm diameter and 5 cm long
Distributed by: | Amersham
Manufactured by:
References
1) | Dosimetry of 192Ir wires for LDR interstitial brachytherapy following the AAPM TG-43 dosimetric formalism
Med Phys 28(2001),156-166
P. Karaiskos, P. Papagiannis, A. Angelopoulos, L.Sakelliou, D. Baltas, P. Sandilos, L. Vlachos
2) | “Dosimetric characterization of Ir-192 LDR elongated sources”
Med Phys 35(2008)1154
R. van der Laarse, D. Granero, J.Perez-Calatayud, A. S. Meigooni and F. Ballester
Other References
3) | Assessment of the relative dose distribution around an Ir-192 line source using a liquid ionization chamber
Med. Phys. 26, 1932-1941 (1999)
J Bahar-Gogani et al
4) | "Monte Carlo calculation of dose rate distributions around 0,5 and 0,6 mm in diameter 192Ir wires"
Med Phys 26(1999)395-401
F. Ballester, C. Hernandez,V. Carmona, J. Pérez-Calatayud and F. Lliso
5) | "Monte Carlo calculation of dose rate distributions around 192Ir wires"
Med Phys 24(1997)1221-1228
F. Ballester, C. Hernandez, J. Pérez-Calatayud and F. Lliso
Calculation details
Method: | Monte Carlo
Code:
Version:
Phantom material: | Water
Phantom size: | Sphere 30 cm in diameter
Cutoff energy: | 10 keV
Cross-section:
Electron tracking: | No
Revised
2006-01-31

## Source design and materials
Source design and materials | Ir-192 LDR wire 0.3 mm diameter and 5 cm long
The central core is 0.1 mm in diameter, composed of about 25% Ir and 75% Pt, encased in a 0.1 mm Pt sheath
The Ir-192 wires were modeled as cylinders: an inner cylindrical core 0.1 mm in diameter encased in a 0.1 mm Pt cylindrical sheath with the same heigh

## Geometry Function
Geometry Function |  |  |  | Ir-192 LDR wire 0.3 mm diameter and 5 cm long
Geometry factor G(r,q) for L = 5 cm.
To use these data, see table I and Section II.B in Karaiskos et al.
Radial distance r (cm)
Angle (º) | 0.1 | 0.2 | 0.3 | 0.4 | 0.5 | 0.6 | 0.7 | 0.8 | 0.9 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 7 | 10
1 | 205.4 | 179.9 | 119.8 | 89.8 | 71.8 | 59.8 | 51.3 | 44.8 | 39.8 | 35.811 | 23.751 | 17.557 | 7.2 | 0.363 | 0.103 | 0.053 | 0.034 | 0.023 | 0.011
2 | 102.5 | 89.9 | 59.9 | 44.8 | 35.8 | 29.8 | 25.6 | 22.3 | 19.8 | 17.813 | 11.753 | 8.561 | 3.601 | 0.359 | 0.102 | 0.053 | 0.034 | 0.023 | 0.011
3 | 68.3 | 59.9 | 39.9 | 29.9 | 23.8 | 19.8 | 17 | 14.8 | 13.2 | 11.815 | 7.754 | 5.566 | 2.401 | 0.354 | 0.102 | 0.053 | 0.034 | 0.023 | 0.011
4 | 51.1 | 44.9 | 29.9 | 22.4 | 17.9 | 14.8 | 12.7 | 11.1 | 9.83 | 8.817 | 5.756 | 4.073 | 1.801 | 0.347 | 0.102 | 0.053 | 0.034 | 0.023 | 0.011
5 | 40.9 | 35.9 | 23.9 | 17.9 | 14.3 | 11.6 | 10.1 | 8.833 | 7.83 | 7.019 | 4.558 | 3.181 | 1.442 | 0.339 | 0.101 | 0.053 | 0.034 | 0.023 | 0.011
7 | 29.1 | 25.6 | 17 | 12.7 | 10.2 | 8.42 | 7.19 | 6.267 | 5.53 | 4.966 | 3.191 | 2.171 | 1.031 | 0.321 | 0.101 | 0.053 | 0.033 | 0.023 | 0.011
10 | 20.3 | 17.9 | 11.9 | 8.88 | 7.07 | 5.86 | 5 | 4.346 | 3.84 | 3.43 | 2.171 | 1.431 | 0.724 | 0.291 | 0.099 | 0.052 | 0.033 | 0.023 | 0.011
15 | 13.5 | 12 | 7.93 | 5.91 | 4.69 | 3.88 | 3.3 | 2.858 | 2.52 | 2.241 | 1.386 | 0.883 | 0.486 | 0.246 | 0.094 | 0.051 | 0.033 | 0.023 | 0.011
20 | 10.2 | 9.03 | 5.96 | 4.43 | 3.51 | 2.89 | 2.45 | 2.122 | 1.86 | 1.653 | 1.002 | 0.629 | 0.367 | 0.21 | 0.089 | 0.05 | 0.032 | 0.023 | 0.011
25 | 8.16 | 7.27 | 4.79 | 3.55 | 2.81 | 2.31 | 1.95 | 1.685 | 1.48 | 1.306 | 0.78 | 0.487 | 0.297 | 0.183 | 0.085 | 0.049 | 0.032 | 0.023 | 0.011
30 | 6.85 | 6.12 | 4.03 | 2.98 | 2.35 | 1.93 | 1.63 | 1.4 | 1.22 | 1.079 | 0.637 | 0.399 | 0.251 | 0.163 | 0.08 | 0.047 | 0.031 | 0.023 | 0.01
35 | 5.93 | 5.32 | 3.49 | 2.58 | 2.03 | 1.66 | 1.4 | 1.2 | 1.05 | 0.921 | 0.539 | 0.339 | 0.219 | 0.147 | 0.076 | 0.046 | 0.03 | 0.022 | 0.01
40 | 5.26 | 4.73 | 3.1 | 2.28 | 1.79 | 1.47 | 1.23 | 1.055 | 0.917 | 0.807 | 0.469 | 0.297 | 0.195 | 0.134 | 0.072 | 0.044 | 0.03 | 0.022 | 0.01
45 | 4.75 | 4.28 | 2.8 | 2.06 | 1.62 | 1.32 | 1.11 | 0.946 | 0.821 | 0.721 | 0.418 | 0.266 | 0.178 | 0.124 | 0.068 | 0.043 | 0.029 | 0.022 | 0.01
50 | 4.36 | 3.94 | 2.57 | 1.89 | 1.48 | 1.2 | 1.01 | 0.862 | 0.747 | 0.656 | 0.379 | 0.242 | 0.164 | 0.116 | 0.066 | 0.042 | 0.029 | 0.021 | 0.01
60 | 3.82 | 3.47 | 2.26 | 1.65 | 1.29 | 1.05 | 0.88 | 0.747 | 0.647 | 0.566 | 0.327 | 0.211 | 0.145 | 0.105 | 0.061 | 0.04 | 0.028 | 0.021 | 0.01
70 | 3.5 | 3.18 | 2.07 | 1.51 | 1.18 | 0.96 | 0.8 | 0.679 | 0.587 | 0.513 | 0.296 | 0.192 | 0.134 | 0.098 | 0.058 | 0.038 | 0.027 | 0.02 | 0.01

## Dose Rate Constant
Dose-Rate Constant |  |  | Ir-192 LDR wire 0.3 mm diameter and 5 cm long
L = | 0.521 | cGy/(h·U)

## Radial Dose Function
Radial-Dose Function Table |  | Ir-192 LDR wire 0.3 mm diameter and 5 cm long
r (cm) | g( r )
0.2 | 0.949
0.3 | 0.964
0.5 | 0.98
0.7 | 0.989
1 | 1
1.5 | 1.011
2 | 1.017
3 | 1.02
4 | 1.017
5 | 1.01
6 | 0.997
7 | 0.981
8 | 0.956
9 | 0.93
10 | 0.897

## Anisotropy Function
Anisotropy Function Table | [5 empty] | Ir-192 LDR wire 0.3 mm diameter and 5 cm long
F(r,q)
r (cm)
q (º) | 0.5 | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 10
0.5 |  |  |  | 0.526 | 0.355 | 0.378 | 0.418 | 0.462 | 0.488 | 0.547
2.5 |  |  |  | 0.829 | 0.577 | 0.547 | 0.561 | 0.577 | 0.595 | 0.643
5.5 |  |  | 0.941 | 0.901 | 0.752 | 0.71 | 0.709 | 0.717 | 0.731 | 0.756
7.5 |  |  | 0.944 | 0.945 | 0.918 | 0.813 | 0.778 | 0.773 | 0.783 | 0.789
10.5 |  |  | 0.953 | 0.956 | 0.937 | 0.867 | 0.84 | 0.838 | 0.835 | 0.836
15.5 | 0.963 | 0.964 | 0.965 | 0.956 | 0.917 | 0.899 | 0.895 | 0.893 | 0.893 | 0.899
20.5 | 0.976 | 0.968 | 0.972 | 0.966 | 0.942 | 0.931 | 0.927 | 0.924 | 0.925 | 0.93
25.5 | 0.985 | 0.976 | 0.976 | 0.973 | 0.958 | 0.951 | 0.95 | 0.95 | 0.95 | 0.952
30.5 | 0.99 | 0.98 | 0.983 | 0.977 | 0.968 | 0.963 | 0.963 | 0.961 | 0.963 | 0.961
35.5 | 0.991 | 0.985 | 0.985 | 0.982 | 0.978 | 0.972 | 0.972 | 0.973 | 0.973 | 0.973
40.5 | 0.992 | 0.986 | 0.987 | 0.986 | 0.983 | 0.981 | 0.978 | 0.977 | 0.98 | 0.978
45.5 | 0.991 | 0.989 | 0.991 | 0.99 | 0.985 | 0.984 | 0.984 | 0.984 | 0.983 | 0.984
50.5 | 0.992 | 0.991 | 0.992 | 0.992 | 0.99 | 0.99 | 0.99 | 0.987 | 0.989 | 0.989
55.5 | 0.993 | 0.992 | 0.996 | 0.995 | 0.994 | 0.992 | 0.993 | 0.992 | 0.99 | 0.989
60.5 | 0.994 | 0.993 | 0.996 | 0.997 | 0.997 | 0.994 | 0.994 | 0.993 | 0.995 | 0.996
65.5 | 0.994 | 0.996 | 0.998 | 0.997 | 0.999 | 0.997 | 0.998 | 0.997 | 0.998 | 0.998
70.5 | 0.995 | 0.992 | 0.998 | 0.999 | 0.997 | 0.998 | 0.999 | 0.997 | 0.998 | 0.998
75.5 | 0.998 | 0.996 | 0.999 | 1001 | 0.999 | 0.998 | 0.999 | 0.997 | 0.998 | 0.995
80.5 | 1 | 0.998 | 0.999 | 1 | 1001 | 0.998 | 0.999 | 1 | 0.998 | 0.999
85.5 | 1 | 1 | 0.998 | 1 | 1 | 0.999 | 0.999 | 0.998 | 0.999 | 0.99
90 | bar: 0.5 to r (cm) 10

## D(y,z)
Absolute dose rate table: D(y,z) (cGy/h U)
Ir-192 LDR wire 0.3 mm diameter and 5 cm long
This data have been taken from van der Laarse et al
Away y (cm)
Along | 0 | 0.25 | 0.5 | 0.75 | 1 | 1.5 | 2 | 2.5 | 3 | 4 | 5 | 6 | 8 | 10 | 14
z (cm)
0 | ---- | 2.47 | 1177 | 0.741 | 0.522 | 0.304 | 0.1998 | 0.1407 | 0.104 | 0.0627 | 0.0415 | 0.0292 | 0.01645 | 0.01029 | 0.0048
0.25 | ---- | 2.47 | 1176 | 0.74 | 0.52 | 0.303 | 0.1992 | 0.1402 | 0.1036 | 0.0625 | 0.0414 | 0.0292 | 0.01644 | 0.01028 | 0.0048
0.5 | ---- | 2.46 | 1172 | 0.735 | 0.516 | 0.301 | 0.1972 | 0.1388 | 0.1026 | 0.0621 | 0.0412 | 0.0291 | 0.01638 | 0.01026 | 0.0048
0.75 | ---- | 2.46 | 1164 | 0.727 | 0.509 | 0.295 | 0.1937 | 0.1365 | 0.1011 | 0.0614 | 0.0408 | 0.0289 | 0.01632 | 0.01023 | 0.00478
1 | ---- | 2.44 | 1150 | 0.715 | 0.499 | 0.288 | 0.1887 | 0.1334 | 0.099 | 0.0604 | 0.0403 | 0.0286 | 0.01621 | 0.01018 | 0.00477
1.5 | ---- | 2.39 | 1098 | 0.671 | 0.462 | 0.265 | 0.1747 | 0.1243 | 0.0931 | 0.0577 | 0.0389 | 0.0278 | 0.01591 | 0.01006 | 0.00474
2 | ---- | 2.22 | 0.966 | 0.576 | 0.396 | 0.231 | 0.1548 | 0.1122 | 0.0853 | 0.054 | 0.0371 | 0.0268 | 0.01552 | 0.00989 | 0.00469
2.5 | ---- | 1259 | 0.621 | 0.405 | 0.296 | 0.1864 | 0.1311 | 0.098 | 0.0762 | 0.0499 | 0.0349 | 0.0255 | 0.01505 | 0.00966 | 0.00463
3 | 0.1709 | 0.296 | 0.276 | 0.234 | 0.196 | 0.1415 | 0.1067 | 0.0833 | 0.0668 | 0.0453 | 0.0324 | 0.0241 | 0.0145 | 0.0094 | 0.00455
4 | 0.0427 | 0.0724 | 0.0863 | 0.0891 | 0.0871 | 0.0779 | 0.0672 | 0.0574 | 0.049 | 0.0362 | 0.0273 | 0.0211 | 0.01324 | 0.00881 | 0.00437
5 | 0.0241 | 0.0346 | 0.0427 | 0.0461 | 0.0473 | 0.0462 | 0.0431 | 0.0393 | 0.0353 | 0.0282 | 0.0224 | 0.01799 | 0.01186 | 0.00812 | 0.00416
6 | 0.01669 | 0.0211 | 0.0258 | 0.0282 | 0.0295 | 0.0301 | 0.0292 | 0.0277 | 0.0258 | 0.0219 | 0.01822 | 0.01513 | 0.01048 | 0.0074 | 0.00392
8 | 0.00969 | 0.01093 | 0.01273 | 0.01389 | 0.01462 | 0.01535 | 0.01542 | 0.0152 | 0.01475 | 0.01347 | 0.01202 | 0.01055 | 0.00796 | 0.00597 | 0.0034
10 | 0.00643 | 0.00688 | 0.00764 | 0.00824 | 0.00866 | 0.0091 | 0.00929 | 0.00929 | 0.0092 | 0.00874 | 0.0081 | 0.00739 | 0.00596 | 0.0047 | 0.00288
14 | 0.00334 | 0.00343 | 0.0036 | 0.00378 | 0.00391 | 0.00409 | 0.0042 | 0.00423 | 0.00424 | 0.00418 | 0.00402 | 0.00383 | 0.00336 | 0.00285 | 0.00196