## General Information
| General Information | Pd-103 - Best Medical model 2335 source
Source Information
Source name: | Best Medical model 2335 source
Radionuclide: | Pd-103
Model Number: | 2335 source
Distributed by:
Manufactured by:
References
1) | "Supplement to the 2004 update of the AAPM Task Group Nº43 Report"
Med Phys 34(6), June 2007(2187-2205)
M.J. Rivard, W.M.Butler, L.A. DeWerd, M.S. Huq, G.S Ibbott, A.S. Meigooni, C.S. Melhus, M G. Mitch, R. Nath, J. F. Williamson
Calculation details
Method:
Code:
Version:
Phantom material:
Phantom size:
Cutoff energy:
Cross-section:
Electron tracking:
Revised

## Geometry and materials
Geometry and materials | Pd-103 - Best Medical model 2335 source

## Geometry Function
Geometry Function |  |  | Pd-103 - Best Medical model 2335 source
Source length for Line source approximation:
L= | 4.55 | mm

## Dose-Rate Constant
Dose-Rate Constant |  |  | Pd-103 - Best Medical model 2335 source
L = | 0.685 | cGy/(hU)

## Anisotropy Function Table
Anisotropy Function Table |  |  |  | Pd-103 - Best Medical model 2335 source
F(r,q)
r (cm)
q (º) | 1 | 2 | 3 | 4 | 5 | 6 | 7
0 | 0.797 | 0.69 | 0.674 | 0.672 | 0.663 | 0.675 | 0.63
5 | 0.801 | 0.696 | 0.683 | 0.669 | 0.666 | 0.679 | 0.645
10 | 0.79 | 0.69 | 0.673 | 0.675 | 0.665 | 0.69 | 0.644
15 | 0.675 | 0.613 | 0.608 | 0.604 | 0.626 | 0.62 | 0.581
20 | 0.608 | 0.591 | 0.596 | 0.601 | 0.616 | 0.647 | 0.595
25 | 0.675 | 0.639 | 0.637 | 0.659 | 0.653 | 0.706 | 0.651
30 | 0.681 | 0.66 | 0.679 | 0.694 | 0.694 | 0.717 | 0.672
35 | 0.725 | 0.693 | 0.705 | 0.721 | 0.703 | 0.73 | 0.687
40 | 0.762 | 0.736 | 0.75 | 0.747 | 0.741 | 0.775 | 0.72
45 | 0.792 | 0.807 | 0.846 | 0.847 | 0.866 | 0.876 | 0.804
50 | 0.885 | 0.88 | 0.885 | 0.887 | 0.909 | 0.907 | 0.835
60 | 0.915 | 0.929 | 0.944 | 0.936 | 0.965 | 1.001 | 0.912
70 | 0.932 | 0.96 | 0.972 | 0.965 | 0.975 | 1.014 | 0.916
80 | 0.941 | 0.975 | 0.986 | 0.985 | 0.999 | 1.017 | 0.915
90 | 1 | 1 | 1 | 1 | 1 | 1 | 1

## Radial-Dose Function Table
Radial-Dose Function Table |  |  | Pd-103 - Best Medical model 2335 source
Line source approximation,
L = | 4.55 | mm | Point source approximation
r (cm) | gL(r) |  | r (cm) | gp(r)
0.1 | 0.826 |  | 0.1 | 0.427
0.15 | 1.066 |  | 0.15 | 0.706
0.25 | 1.236 |  | 0.25 | 1.02
0.5 | 1.307 |  | 0.5 | 1.247
0.75 | 1.128 |  | 0.75 | 1.114
1 | 1 |  | 1 | 1
1.5 | 0.742 |  | 1.5 | 0.749
2 | 0.533 |  | 2 | 0.539
3 | 0.296 |  | 3 | 0.3
4 | 0.158 |  | 4 | 0.161
5 | 0.092 |  | 5 | 0.0935
6 | 0.0529 |  | 6 | 0.0538
7 | 0.0309 |  | 7 | 0.0314
8 | 0.018 |  | 8 | 0.0184
9 | 0.0105 |  | 9 | 0.0107
10 | 0.0062 |  | 10 | 0.0063

## Anisotropy Factor
Anisotropy Factor |  |  | Pd-103 - Best Medical model 2335 source
| r (cm) | fan(r)
| 0.1 | 1.052
| 0.15 | 1.205
| 0.25 | 1.213
| 0.5 | 0.938
| 0.75 | 0.894
| 1 | 0.879
| 2 | 0.872
| 3 | 0.881
| 4 | 0.881
| 5 | 0.89
| 6 | 0.909
| 7 | 0.845

## QA
Dose rates for QA |  |  | Pd-103 - Best Medical model 2335 source
Transverse plane dose rates (cGy·h−1·U−1) as a function of distance using the 1-D formalism of Eq. (11) with interpolation for gL(r) and Øan(r). |  |  |  | QA along-away table obtained using TG-43U1S1 formalism and consensus dataset
D(y,z) / cGy h-1 U-1
y / cm
| r (cm) | cGy h-1 U-1 | z / cm | 0 | 1 | 2 | 3 | 4 | 5 | 6 | 7
| 0.1 | 30.8 | 0 | --- | 0.685 | 0.092433 | 0.022869 | 0.0068721 | 0.0025619 | 0.0010232 | 0.00043917
| 0.15 | 25.9 | 0.5 | 3.9063 | 0.48166 | 0.082144 | 0.021537 | 0.0066096 | 0.002509 | 0.0010145 | 0.00041831
| 0.25 | 12.6 | 1 | 0.58554 | 0.22047 | 0.062632 | 0.018531 | 0.0060049 | 0.0023551 | 0.00096914 | 0.00038631
| 0.5 | 3.21 | 1.5 | 0.17483 | 0.091897 | 0.041604 | 0.014548 | 0.0051639 | 0.0021097 | 0.00088674 | 0.00035251
| 0.75 | 1.21 | 2 | 0.064892 | 0.04331 | 0.024659 | 0.010472 | 0.0042256 | 0.0018257 | 0.00078279 | 0.00031831
| 1 | 0.602 | 2.5 | 0.031773 | 0.021748 | 0.013491 | 0.0069439 | 0.0032784 | 0.0015214 | 0.00066506 | 0.00027875
| 1.5 | 0.2 | 3 | 0.015532 | 0.01149 | 0.0080757 | 0.0046843 | 0.0023761 | 0.0012029 | 0.00054319 | 0.00023605
| 2 | 0.0806 | 3.5 | 0.008725 | 0.0066212 | 0.0046377 | 0.0029263 | 0.0017674 | 0.00089563 | 0.00042376 | 0.00019275
| 3 | 0.0201 | 4 | 0.004638 | 0.0038164 | 0.002957 | 0.0018407 | 0.0012605 | 0.00066577 | 0.00032698 | 0.00015359
| 4 | 0.00605 | 4.5 | 0.0028778 | 0.0024874 | 0.0018014 | 0.0012894 | 0.00081355 | 0.00047882 | 0.00024848 | 0.00012223
| 5 | 0.00228 | 5 | 0.0017033 | 0.0015531 | 0.0011926 | 0.00087329 | 0.00055552 | 0.00033386 | 0.00018169 | 9.4168e-05
| 6 | 0.00093 | 5.5 | 0.0011182 | 0.0010408 | 0.00076758 | 0.00058643 | 0.00037377 | 0.00023342 | 0.0001315 | 6.975e-05
| 7 | 0.000371 | 6 | 0.00069199 | 0.0006598 | 0.00049707 | 0.00038899 | 0.00025454 | 0.00015758 | 9.6879e-05 | 5.2597e-05
| 8 | 0.000166 | 6.5 | 0.00045136 | 0.00042965 | 0.00031694 | 0.00025109 | 0.00017658 | 0.00011025 | 6.7156e-05 | 3.9292e-05
| 9 | 7.66e-05 | 7 | 0.00027707 | 0.00026823 | 0.00020535 | 0.0001649 | 0.00011706 | 7.8673e-05 | 4.7278e-05 | 2.8068e-05
| 10 | 3.62e-05