## 1031L
| TG-43U1S2 CONSENSUS |  |  |  | IBt Inc. Pd-103 model 1031L
Supplement 2 for the 2004 update of the AAPM Task Group No. 43 Report: Joint recommendations by the AAPM and GEC-ESTRO
Rivard et al.
Med. Phys. 44 (9) (September 2017) e297-e338 http://dx.doi.org/10.1002/mp.12430
Erratum Med. Phys. 45(2) (February 2018) 971-974 http://dx.doi.org/10.1002/mp.12728
| L = | 0.701 | cGy/(h U)
| u(L) = | 0.02 | cGy/(h U)
| L = | 0.435 | cm
| Interpolated data are highlighted in boldface while extrapolated data are underlined. | [18 empty] | D(r,q)
cGy h-1 U-1
[7 empty] | F(r,q) | Distance from Active Source Center, r (cm) | [8 empty] | r | 1D dose (P) |  |  |  | r / cm
| r (cm) | gL(r) L = 0.435 cm | gP(r ) | r (cm) | fan(r) | q / deg | 0.1 | 0.25 | 0.5 | 0.75 | 1 | 2 | 3 | 5 | 7.5 | 10 | cm | cGy h-1 U-1 | q (º) | 0.1 | 0.25 | 0.5 | 0.75 | 1 | 2 | 3 | 5 | 7.5 | 10
| 0.1 | 1.016 | 0.541 | 0.1 | 1.132 | 0 |  | 0.623 | 0.194 | 0.224 | 0.276 | 0.406 | 0.46 | 0.52 | 0.574 | 0.661 | 0.1 | 42.9 | 0 |  | 38 | 0.86 | 0.352 | 0.206 | 0.0421 | 0.01156 | 0.001382 | 0.0001454 | 2.55e-05
| 0.15 | 1.17 | 0.792 | 0.15 | 1.704 | 2 |  | 0.612 | 0.233 | 0.284 | 0.344 | 0.414 | 0.441 | 0.473 | 0.531 | 0.608 | 0.15 | 42 | 2 |  | 36.5 | 1.034 | 0.445 | 0.257 | 0.0429 | 0.01107 | 0.001256 | 0.0001345 | 2.35e-05
| 0.25 | 1.301 | 1.087 | 0.25 | 1.363 | 5 |  | 0.558 | 0.399 | 0.393 | 0.409 | 0.444 | 0.47 | 0.501 | 0.546 | 0.635 | 0.25 | 16.61 | 5 |  | 30.4 | 1.765 | 0.615 | 0.306 | 0.046 | 0.0118 | 0.001331 | 0.0001383 | 2.46e-05
| 0.5 | 1.264 | 1.211 | 0.5 | 0.986 | 7 |  | 0.555 | 0.503 | 0.473 | 0.466 | 0.48 | 0.5 | 0.527 | 0.572 | 0.639 | 0.5 | 3.35 | 7 |  | 27.8 | 2.22 | 0.74 | 0.348 | 0.0497 | 0.01256 | 0.001399 | 0.0001448 | 2.47e-05
| 0.75 | 1.135 | 1.122 | 0.75 | 0.931 | 10 |  | 0.681 | 0.579 | 0.522 | 0.511 | 0.517 | 0.533 | 0.555 | 0.596 | 0.681 | 0.75 | 1.302 | 10 |  | 29.8 | 2.54 | 0.815 | 0.381 | 0.0535 | 0.01338 | 0.001473 | 0.0001511 | 2.63e-05
| 1 | 1 | 1 | 1 | 0.917 | 15 |  | 0.748 | 0.654 | 0.611 | 0.596 | 0.595 | 0.604 | 0.621 | 0.652 | 0.714 | 1 | 0.643 | 15 |  | 26.5 | 2.83 | 0.949 | 0.443 | 0.0616 | 0.01517 | 0.001648 | 0.000165 | 2.76e-05
| 1.5 | 0.766 | 0.772 | 1.5 | 0.908 | 20 |  | 0.917 | 0.696 | 0.674 | 0.67 | 0.662 | 0.669 | 0.675 | 0.708 | 0.734 | 1.5 | 0.219 | 20 |  | 27 | 2.96 | 1.041 | 0.496 | 0.0685 | 0.0168 | 0.001792 | 0.0001793 | 2.84e-05
| 2 | 0.576 | 0.582 | 2 | 0.899 | 25 | 0.896 | 1.027 | 0.739 | 0.716 | 0.713 | 0.712 | 0.719 | 0.728 | 0.748 | 0.794 | 2 | 0.0917 | 25 | 94.6 | 25.9 | 3.08 | 1.097 | 0.526 | 0.0736 | 0.01805 | 0.001931 | 0.0001893 | 3.07e-05
| 3 | 0.316 | 0.32 | 3 | 0.9 | 30 | 0.952 | 1.083 | 0.782 | 0.758 | 0.756 | 0.753 | 0.757 | 0.765 | 0.777 | 0.817 | 3 | 0.0225 | 30 | 82.7 | 23.9 | 3.19 | 1.15 | 0.555 | 0.0777 | 0.01897 | 0.00203 | 0.0001968 | 3.16e-05
| 4 | 0.1718 | 0.1743 | 4 | 0.898 | 40 | 0.897 | 1.1 | 0.869 | 0.835 | 0.831 | 0.828 | 0.832 | 0.838 | 0.851 | 0.873 | 4 | 0.00686 | 40 | 58 | 19.77 | 3.38 | 1.244 | 0.604 | 0.0853 | 0.0208 | 0.00222 | 0.000215 | 3.38e-05
| 5 | 0.0931 | 0.0945 | 5 | 0.898 | 50 | 0.902 | 1.074 | 0.933 | 0.906 | 0.902 | 0.893 | 0.897 | 0.895 | 0.906 | 0.909 | 5 | 0.00238 | 50 | 47.3 | 16.59 | 3.47 | 1.322 | 0.648 | 0.0917 | 0.0225 | 0.00237 | 0.000229 | 3.51e-05
| 6 | 0.0489 | 0.0497 | 6 | 0.902 | 60 | 0.926 | 1.043 | 0.972 | 0.958 | 0.955 | 0.945 | 0.948 | 0.94 | 0.945 | 0.949 | 6 | 0.000872 | 60 | 41.8 | 14.46 | 3.48 | 1.373 | 0.68 | 0.0968 | 0.0237 | 0.00249 | 0.000239 | 3.67e-05
| 7 | 0.0268 | 0.0273 | 7 | 0.905 | 70 | 0.964 | 1.022 | 0.991 | 0.983 | 0.985 | 0.976 | 0.978 | 0.972 | 0.98 | 0.963 | 7 | 0.000353 | 70 | 39.4 | 13.17 | 3.44 | 1.391 | 0.695 | 0.0997 | 0.0244 | 0.00258 | 0.000248 | 3.72e-05
| 8 | 0.01507 | 0.0153 | 8 | 0.909 | 80 | 0.992 | 1.007 | 0.998 | 0.995 | 0.998 | 0.995 | 0.994 | 0.991 | 0.989 | 0.988 | 8 | 0.0001523 | 80 | 38.3 | 12.45 | 3.41 | 1.395 | 0.701 | 0.1015 | 0.0248 | 0.00262 | 0.00025 | 3.82e-05
| 9 | 0.00892 | 0.00905 | 9 | 0.913 | 90 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 9 | 7.15e-05 | 90 | 37.9 | 12.19 | 3.4 | 1.398 | 0.701 | 0.102 | 0.0249 | 0.00265 | 0.000253 | 3.86e-05
| 10 | 0.00543 | 0.00551 | 10 | 0.916 | [11 empty] | 10 | 3.54e-05
D(y,z)
cGy h-1 U-1
y / cm
[19 empty] | z/cm | 0 | 0.5 | 1 | 1.5 | 2 | 3 | 4 | 5 | 6 | 7
[19 empty] | 0 |  | 3.395 | 0.701 | 0.241 | 0.102 | 0.0249 | 0.00763 | 0.00265 | 0.000967 | 0.00039
[19 empty] | 0.5 | 0.86 | 1.484 | 0.514 | 0.204 | 0.0914 | 0.0235 | 0.00733 | 0.00257 | 0.000944 | 0.000382
[19 empty] | 1 | 0.206 | 0.399 | 0.248 | 0.131 | 0.068 | 0.01995 | 0.00657 | 0.00236 | 0.000887 | 0.000364
[19 empty] | 1.5 | 0.0844 | 0.1362 | 0.1112 | 0.0732 | 0.0443 | 0.01541 | 0.00551 | 0.00206 | 0.000802 | 0.000336
[19 empty] | 2 | 0.0421 | 0.0545 | 0.052 | 0.0393 | 0.027 | 0.0111 | 0.00436 | 0.001725 | 0.000701 | 0.000301
[19 empty] | 2.5 | 0.0212 | 0.0249 | 0.0257 | 0.0213 | 0.016 | 0.00763 | 0.0033 | 0.001381 | 0.000591 | 0.000264
[19 empty] | 3 | 0.01156 | 0.01256 | 0.01329 | 0.01184 | 0.00943 | 0.00507 | 0.00241 | 0.001071 | 0.000483 | 0.000223
[19 empty] | 3.5 | 0.00646 | 0.00676 | 0.00716 | 0.00669 | 0.00561 | 0.00333 | 0.001705 | 0.000809 | 0.000386 | 0.0001852
[19 empty] | 4 | 0.00376 | 0.00381 | 0.00401 | 0.00386 | 0.00338 | 0.00216 | 0.001181 | 0.000605 | 0.000302 | 0.0001512
[19 empty] | 4.5 | 0.00225 | 0.00221 | 0.00232 | 0.00227 | 0.00205 | 0.001384 | 0.000809 | 0.000444 | 0.000234 | 0.0001224
[19 empty] | 5 | 0.001382 | 0.001322 | 0.001374 | 0.001357 | 0.001247 | 0.000887 | 0.000557 | 0.000322 | 0.000178 | 9.77e-05
[19 empty] | 5.5 | 0.000845 | 0.000798 | 0.000826 | 0.000822 | 0.000768 | 0.000577 | 0.000381 | 0.000233 | 0.0001346 | 7.71e-05
[19 empty] | 6 | 0.000525 | 0.000492 | 0.00051 | 0.000509 | 0.000484 | 0.00038 | 0.000261 | 0.0001673 | 0.000102 | 6.06e-05
[19 empty] | 6.5 | 0.000338 | 0.000315 | 0.000325 | 0.000324 | 0.000311 | 0.000253 | 0.0001803 | 0.0001207 | 7.7e-05 | 4.74e-05
[19 empty] | 7 | 0.00022 | 0.000205 | 0.000211 | 0.000209 | 0.000202 | 0.0001698 | 0.0001254 | 8.79e-05 | 5.82e-05 | 3.69e-05